 [image: image2.jpg](>

SCIENTZ



         宁波新芝生物科技股份有限公司

                         Ningbo Scientz Biotechnology Co.,Ltd

Scientz-500F Freeze dryer Technical data 

1. General

	Item
	Technical data

	Total Square
	5.4m2

	Ability to capture water
	75 kg/batch 6.5KW

	 Power（50Hz/380V）
	29KW total power not working power  Consume power＜＜29KWh）

	Cooling water
	About 16m3/h（≤25℃）

	Host  L×W×H
	About 4800mm*1600mm*2700mm（Against real drawing ）

	structure
	Whole machine one piece

	Total weight
	About 4500kg

	Compressed air
	0.5 ～ 0.6 MPa

	Power
	3 Phase 5 lines  380V 50Hz


2. Chamber  

	Item
	Technical Data

	Inner chamber material 
	AISI 304
AISI 304 stainless steel 

	Lock method 
	Manual rotation door lock

	Chamber door sealing 
	Fixed type single way door sealing 

	Chamber door open position
	From right to Left (Against client’s requirement )

	Chamber door material 
	AISI 304 Stainless steel 

	Chamber roughness 
	Ra ≤ 0.5 μm

	Chamber shape 
	Rectangle 

	Heating preservation material 
	PU integral eps（Suit with environmental requirement ）

	Outer chamber decoration
	AISI 304 stainless steel draw bench plate 


3  Shelves  

	Item
	Technical Data 

	Tray material 
	AISI 316 stainless steel 

	Surface roughness
	Ra ≤ 0.5 μm

	Tray No
	5＋1

	Tray space 
	80mm 

	Tray size (L*W*H)mm
	About W900 mm × D1200 mm ×H20 mm

	Tray refrigerant 

cooling data 
	+20℃～-40℃≤60 Minutes (free loading ) 

	Tray refrigerant 

Heating  data 
	-40℃～+20℃≤60分钟(free loading )



	Tray lowest temperature
	-50℃

	Tray highest temperature 
	+ 60℃

	Heating power 
	7.5 kW  

	Tray smoothly 
	≤ ±1 mm/m

	Protection hurdle 
	Tray ,metal hose and fluid shunt (wedding )fixed directly ,to ensure that the tray layer in the long time no oil leakage and easy for cleaning ,plate layer of left ,right after all three sides protection article in order to prevent the plate /bottle than the plate 

	Tray top press 
	Hydraulic top press 


 Condenser  
	Item 
	Technical  data 

	Condenser material 
	AISI 304 stainless steel 

	Condenser shape 
	Cylinder-shaped

	Coiler material 

	ISI 304 stainless steel tube 

	Coiler shape
	Outer coiler 

	Coiler area
	About 6.6 m2

	Inner surface roughness
	Ra ≤ 0.5 μm

	Condenser cooling data
	+20℃～-40℃≤30min(no loading )

	Condenser lowest Tem
	-70℃

	Defrost mode  

	 electric defrosting 

	Temperature sensor 
	Platinum resistor

	Heating 

preservation material
	PU integral eps（Suit with environmental requirement 

	Outer chamber decoration
	AISI 304 stainless steel drawbench plate 


Refrigeration System  
	Item 
	Technical data 

	Compressor type &Qty 
	S6H-20.2Y(Germany BITZER ,assemble in Beijing) 1set 

	Compressor rating power 
	15 KW

	Refrigerant 
	R507 (International famous low tem ,high energy environment protection refrigerant quality 

	Cool heat exchanger 
	SWEP (Sweden )

	Tray refrigeration
	1 set (indirect refrigeration)

	Cold trap refrigeration
	1set (direct refrigeration)

	Oil separator 
	Fasike (Tianjin)

	Oil plane lens
	DANFOSS (DENMARK)

	Prevent aborb high pressure 

Auto prevent device 
	DANFOSS (DENMARK)

	Refrigerant high pressure 

Controller device 
	DANFOSS (DENMARK)

	Gas liquid separator 
	Fasike |(Tianjin)

	disacidfy filter dryer 
	Fasike |(Tianjin)

	Refrigerant expanding valve 
	DNAFOSS (DENMARK)

	Pressure meter 
	REFCO (Swiss) oil immersed  pressure meter 

	Safety valve 
	Caster (Italy )

	Solenoid valve 

	DANFOSS(DENMARK)


Vacuum System  
	Item
	Technica data 

	(vacuum pump mode 

and Qty)
	Baosi BVS30 1 SET 

	Vacuum pump rating power 
	3.5 KW

	Vacuum sensor 
	Chengdou Zhenghua 

	Cod Chamber valve 

	DN150 Butterfly valve 

	Chamber pump valve 
	Ball valve 

	Final vacuum

	≤ 2  Pa

	air exhaust rate 

	Barometric pressure →10 Pa  ≤ 30 minutes （drying condition）


Cycle System  
	Item 
	Technical data

	Pipeline material

	AISI 304 stainless steel 

	Circulation pump
	Denmark Grundfos wet rotor shield pump US25-120 1 SET

	Circulation rated power
	0.46 KW

	Circulation flow liquid 
	Low viscosity silicone oil 

	Heat exchanger 
	Beijing yaxing zhuoyue Technology Development Co.,Ltd 

	Heating rate 
	Over 1℃/min

	Tem measuring 

head (inter/outer )
	Platinum resistor

	Heating power 
	3 Kw  

	Heating method 
	Solid state relay

	Tray 

highest temperature
	≥ +65℃ (Free loading ,silicone inn

	Fluid 

temperature controller 
	DANFOSS(DENMARK)


8  Contral System 
	Item 
	Technical data

	Controller mode 
	Manual /auto controller 

	PLC
	OMRON

	Touch display 
	OMRON

	Design & assemble standard 
	In conformity with the relevant standards of the state ,public line ,input and output signal wire grouping ,the color difference is obvious ,high ,light voltage low part on the wire color have obvious distinguish ,have connection number 

	Power switch 
	DELIXI

	Intermediate relay
	Japan Hequan

	A C contactor
	Schneider

	Tem measure probe 
	PT-100 platinum resistor


11  Protect System  
	Item 
	Technical data 

	Water pressure alarm 
	Yes ,water flow relay

	Hydraulic 

difference protection
	Bitzer freezer origin binding 

	High pressure protection
	DANFOSS (DENMARK)

	Electronic 

heating protection
	BITZER freezer original built-in 

	Automatic liquid function
	Yes DANFOSS (DANMARK)

	Over temperature alarm 
	YES

	Circulator pressure alarm 
	YES


 pictures

[image: image1.jpg]



1
5

[image: image2.jpg]